## Slide 1
Characteristics of the Irradiated Hamamatsu p-Bulk Silicon Microstrip Sensors
K. Hara1, H. Hatano1, T. Meguro1, S. Mitsui1, T. Okuyama1, M. Yamada1,
Y. Ikegami2, T. Kohriki2, S. Terada2, Y. Unno2,
K. Yamamura3, and S. Kamata3
1IPAS, University of Tsukuba
2INPS, KEK
3Solid State Div., Hamamatsu Photonics
IEEE NSS Dresden, N19-2 (Radiation Damage Effects on Detection Materials II)
Notes: <number>

## Slide 2
We are developing non-inverting n+-on-p sensors
lack of experience with sensor-quality p-type Si (wafer availability)‏
full depletion voltage & collectable charge
reliable strip isolation
punch through (for sensor protection)‏
annealing property
(Lorentz angle – measurement started)‏
….
Evaluate these of irradiated sensors
made with p-wafers available to HPK
Development of p-Bulk Sensors
Microstrip sensors at SuperLHC will receive 1015 n/cm2
n-type FZ Si inverts into p-type after < 1014 n/cm2
(accumulates donor type impurities by irradiation)‏
p+-on-n (present ATLAS microstrip) p+-on-p: need fully depleted
n+-on-n (present ATLAS pixel n+-on-p ): need double-side process
Depletion develops from strip side.
Still work at partial depletion, if VFD exceeds the applicable voltage *
P-bulk stays P
n+
p+
depletion
n+-on-p sensor
*VMAX=500 V to re-use ATLAS system

## Slide 3
Evaluated Samples
Float Zone: FZ – 6”, FZp – 6”, FZ – 4”
Magnetic Cz: MCZ - 6”, MCZ – 4”
FZp has fewer defects than standard FZ
wafers
strip isolation
P-stop … with different widths & three doping densities
P-spray … added to low density P-stop, or P-spray only
no P-stop, no P-spray
miniatures : 10 mm x 10 mm, with 100 strips at ~80 m pitch
test structure : evaluate punch-through properties
punch-through design (distance between n+ electrode and biasring )‏
In case of beam splash causing many charges created, the voltage across AC oxide may exceed its rating (~100V for HPK).
create a route to discharge by punch-through mechanism

## Slide 4: Parameters in Detail
ATLAS07 HPK6”
<100>
5
<100>
2
<100>
5
ATLAS06 HPK6”
n/a
<100>
0.8
<111>
6
ATLAS05 HPK4”*
FZp
MCZ
FZ
process
8x1012
( 2x1012 )‏
10x1012
20x1012
2x1012
( 2x1012 )‏
5x1012
20x1012
P-stop
( P-spray )‏
P-stop
LO/Mid
P-stop
HI
doping density (/cm2)‏
crystal orientation &  (kcm)‏
design revised …
LO at 4” is too low
field plates are not efficient
full common is leaky
ATLAS07 is evaluated by a wider collaboration.
We report
narrow common p-stop (~Z3)
no structure (~Z1)
from 1st delivery: modified samples are arriving
(ATLAS05/06)‏
*HPK4” is not available

## Slide 5
Annealing (ATLAS06)
Samples: 6” FZ/FZp/MCZ irradiated with protons to 2x1014 /cm2
VFD evaluated from C-V and laser CC curves
Characteristics similar to n-bulk sensors (min. at 80 min)‏
MCZ shows a larger increase

## Slide 6: Strip Isolation (ATLAS06 results)‏
P-spray (+P-stop)‏
P-stop (mid)‏
P-stop (HI)‏
Fz_pure
Fz_std
MCZ
FZ_pure & p-spray better below 2x1014 (Z1)‏
MCZ better in entire range
Mid & P-spray degrade above 1015 n/cm
Isolation achieved below ~200V bias (<operation voltage) also for Z1
Are these real?
Isolation voltage (Vb@Rinterstrip>10Rbias)

## Slide 7: Strip Isolation (60Co at low rate)‏
Strip isolation primarily determined by surface conditions
Should be testable with 60Co irradiation
biasring
DC pad
A
Isolation data taken while irradiated with 60Co at 100 Gy/h
no structure (Z1)
FZ/MCZ
~100 Gy/h expected at
r =30cm and
L=1035/cm2/s
(44 Gy/h if mips)
ATLAS07

## Slide 8: Strip Isolation (60Co at low rate)‏
Among wafers (Z3, ave of 2 samples each)
Vbias
FZp wafers
FZp better than others, but isolation voltage increases for all.
no structure (Z1) behaves worse
high density p-stop is better than p-spray

## Slide 9
Punch Through - ATLAS06 test structure -
Small difference among wafers (~7V larger for irradiated MCZ)‏
No significant difference between w/ and w/o spray
Fz :Pstop+spray
VPT [V]@R=0.9M
MCZ 400
400
300
200
100
bias
(p-stop)‏
A
8,12,20
“C” with

## Slide 10
PTP Implementation - ATLAS07 -
no structure
narrow common
P-STOP
n+
n+
12
12
Current across a strip to bias-ring at Vbias=300V (bias resistor : ~1.5M
Z1 (PT)‏
Z3 (PT)‏
ALTAS07 irradiated with 70MeV protons
Smaller PT shifts in negative voltage
(beam splash creates negative signal)
Bias-ring
DC pad
A

## Slide 11: PT under 60Co at low rate‏
Similar PT voltages w/ and w/o P-spray after some irradiation
Z3 (PT)‏

## Slide 12
SUMMARY
We are developing non-inverting n+-on-p microstrip sensors, fabricated by Hamamatsu Photonics using commercially available p-type FZ and MCZ wafers.
FZ_pure wafer is selected for ATLAS07.
Strip isolation
  High density 2x1013 /cm2 p-stop preferred for irradiation exceeding 1x1015 /cm2 and ~1kGy/h (beyond sLHC).
  No structure – only p-spray of 2x1012 /cm2 – showed worse isolation than p-stop at 100Gy/h (~sLHC).
  We need to examine these with real size sensors with readout electronics connected.
Punch-through protection is implemented, which turns on around 40 V,
Other characteristics are being evaluated.

## Slide 13
ＢＡＣＫ ＵＰ

## Slide 14: Charge Collection Efficiency
ATLAS05/06

## Slide 15: Proton Irradiation at CYRIC
Cyclotron (70 MeV: e.g. 8h for 1016/cm2 irradiation with 20cm2 scan )‏
BEAM profile @32line
Sample holder
samples
Profile monitor
<number>
Samples cooled to -10oC and biased,
scanned to perform uniform irradiation

## Slide 16
laser head
XY-stage
amps
collimators
focusing lens
Gamp=5mV/fC
Charge Collection Evaluation with Laser
Pulsed Nd:YAG(1064nm) laser can simulate passage of charge particle since it penetrates through Si. The above traces are for Fz samples irradiated to 0.2~5ｘ1015 n/cm2 (Vb=1kV)‏

## Slide 17: Full Depletion Voltage
ATLAS05/06

## Slide 18
ATLAS07 Zone Structure