## Slide 1: STJ developments for FIR photon spectroscopy at Tsukuba
Yuji Takeuchi (Univ. of Tsukuba)
Dec. 18, 2012
SCD review @ KEK 2-Go-kan Bldg.
Contents
Motivation
Hf-STJ development
Nb/Al-STJ response
Development Nb/Al-STJ readout
  FNAL test/SOI opamp/SOI-STJ

## Slide 2: Collaboration Members (Japan-US collab.: Search for Neutrino Decay)
Dec. 2012
Japan Group
Shin-Hong Kim, Yuji Takeuchi, Shinya Kanai, Kazuki Nagata,
Kota Kasahara, Takuya Okudaira (University of Tsukuba) ,
Hirokazu Ikeda, Shuji Matsuura, Takehiko Wada (JAXA/ISAS) ,
Hirokazu Ishino, Atsuko Kibayashi, Yasuki Yuasa(Okayama University) ,
Takuo Yoshida, Yusuke Shimizu, Mikiya Nagashima (Fukui University) ,
Satoshi Mima (RIKEN),
Yukihiro Kato (Kinki University) ,
Masashi Hazumi (KEK)
US Group
Erik Ramberg, Mark Kozlovsky, Paul Rubinov, Dmitri Sergatskov, Jonghee Yoo (Fermilab)
Korea Group
Soo-Bong Kim (Seoul National University)

## Slide 3: Motivation
Expected spectrum for
Sharp edge with
1.9K smearing
Red shift effect
dN/dE(A.U.)
Search for in cosmic neutrino background (CNB)
  Direct detection of CNB
  Direct measurement of neutrino mass
Aiming at sensitivity of detecting from decay for
  Current lower limit
  SM expectation
  L-R symmetric model predicts
SU(2)L ｘ SU(2)R ｘ U(1)

## Slide 4: Detector requirements
Requirements for detector
  Energy measurement for single photon with better than 2% resolution for (, far infrared photon)
  Rocket and satellite experiment with this detector
Superconducting Tunneling Junction (STJ) detectors in development
  Array of 50 Nb/Al-STJ cell with diffraction grating covering
    For rocket experiment aimed at launching in 2016 in earliest, aiming at improvement of lower limit for by 2 order
  STJ using Hafnium: Hf-STJ for satellite experiment (after 2020)
    : Superconducting gap energy for Hafnium
    for 25meV photon: if Fano-factor is less than 0.7 (No back tunneling assumed)

## Slide 5: Dilution refrigerator operation
Our record minimum temperature: 28mK
  4 samples, 1 optical fiber, and RuOx sensor are mounted on the stage
  RuOx sensor is calibrated at known Hf Tc (130mK)
  Goal for Hf-STJ operation: 20mK
28mK
Feb. 2012
Temperature(mK)
100
0
Resistance(Ω)
1
0
2
Hf wire
200
SC transition
Tc=130mK
The dilution refrigerator in this talk is provided by Prof. Ootuka (U. of Tsukuba)

## Slide 6: Hf-STJ development
We succeeded in observation of Josephson current by Hf-HfOx-Hf barrier layer for the first time in the world in 2010.
@T=120-130mK
However, to use this as a detector, much improvement in leak current is required. ( is required to be at pA level or less)
Oxidative condition
1 hour in 10Torr Oxygen ambience
250nm
250nm

## Slide 7: Hf-STJ development
We re-established Hf-STJ process at KEK clean room after 2011 earthquake.
Hf(350nm)
Hf(250nm)
Si wafer
HfOx：20Torr,1hour
anodic oxidation：45nm
B = 0 Gauss
B = 10Gauss
10μV/DIV
50μA/DIV
T=90mK~113mK
Averaged by 128 times on Oscilloscope
However, much worse leakage current than before ( : 5 times higher)

## Slide 8: Hf-STJ development
Since we adopted anodic oxidation, it’s not due to leak from edge region.
Other several samples have similar leakage currents independent of magnitude of .
We suspect this large leakage current is due to noise originated from readout electronics.
  Our study on Nb/Al-STJ leakage current indicates this.
We want to measure I-V curve in low noise environment.

## Slide 9
FIR photon spectroscopy with
diffraction grating + Nb/Al-STJ array
  Diffraction grating covering (16-31meV)
Array of Nb/Al-STJ cell
  We use each Nb/Al-STJ cell as a single-photon counting detector with extremely good S/N for FIR photon of
  for Nb: if consider factor 10 by back-tunneling
Need to develop ultra low temperature (2K) preamplifier
  In collaboration with Fermilab Milli-Kelvin Facility group (Japan-US collaboration: Search for Neutrino Decay)
Nb/Al-STJ array

## Slide 10: Nb/Al-STJ response to laser pulses
Response to DC-like visible laser light (, 2.64eV)
V
I
2Δ
w/ light
w/o light
I-V curve of Nb/Al-STJ sample
laser off
laser on
1MHz,~104photons/pulse
T100mK, Magnet on
Laser ON
ms
4ms/DIV
5μA/DIV
Vbias=1mV
1MHz,~104photons/pulse
Output current from Nb/Al-STJ
Magnet on
The Nb/Al-STJ in this talk is 100umx100um STJ provided by Mima-san (Riken)

## Slide 11: Nb/Al-STJ response to laser pulses
STJ
1kΩ
5pF
Pre-Amp
Shaper
Source meter
V
Response to visible laser pulse (, 2.64eV)
10MHz,30pulses (3us)
Laser

## Slide 12: Nb/Al-STJ leakage current
1mV
250nA
B = 35Gauss
T=1.7K
Leakage: 250nA @1mV
Source Meter
STJ
I_High
I_Low
V_High
V_Low
V: 0.5mV/DIV, I: 500nA/DIV
1mV
10A
B = 35Gauss
T<1.5K
Leakage: 10A @1mV
STJ
1k
I
V
Read
Read
Same sample
Keithley 2636A
Tektronix ADA400A
70Hz
Readout system (noise from readout electronics) affects leakage current!

## Slide 13: Nb/Al-STJ I-V curve measurement at FNAL
Several candidates of HEMT amp will be tested below 2K at FNAL in Dec. 2012 – Feb. 2013.
We will perform a combined test at FNAL in Mar. 2013 to measurement of Nb/Al-STJ I-V curve as well as measurement of Nb/Al-STJ response to visible light at FNAL. two graduated students from Tsukuba will join in this test.
STJ
1k
I
V
Read
Read

## Slide 14: Pre-amplifier developments for Nb/Al-STJ
As a first step, we aim at single visible photon ().
  if consider factor 10 by back-tunneling (Q=1.6fC)
  Requirements for pre-amplifier
    Operate at T=1.8K
    Noise: <0.34fC (2100e)
    Gain: 40mV/fC (64mV for Q=1.6fC)  CF=0.025pF
Useful for Hf-STJ readout as well
STJ
Laser
-
+
-V0
Output
CF

## Slide 15: Test of SOI amplifier at ultra-low temperature
Normal Si JFET is limited by their freeze-out temperature (about 40 K).
We test SOI op amps (FD-SOI-CMOS) developed by Wada-san (JAXA/ISAS) et al., which functions at 4.2K.
This device is provided as Op-amp, the test above is performed at 4.2K by JAXA
Gain=101, Rin=100k, Rf=10M, Input: sine wave f=11Hz

## Slide 16: Test of SOI op amp at Tsukuba
We tested SOI op amp from JAXA at 1.8K.
-
+
Out
10kΩ
1MΩ
In
Gain100
Input 2mV/DIV Output 200mV/DIV
T=1.8K

## Slide 17: Test of SOI op amp at Tsukuba
-
+
r
C
R
Input charge: 150fC (corresponding to input of 100 photons at 470nm into Nb/Al-STJ）
15us
-100mV
Input
1.5mV
2ms
Test at room temperature
Expected signal:
1mV pulse height and 1ms decay for 150fC charge input
We are also in preparation of testing SOI op amp from JAXA as a charge integration amp.
5MΩ
150pF
6.6MΩ

## Slide 18: Development of SOI-STJ
We have started a test of SOI-STJ for our application with Arai-san (KEK).
  Will process STJ layer directly on SOI.
Tests using SOI chips has just started.
  SOI chips with just only pMOS and nMOS FET.
SOI
STJ
Nb
metal pad
G
S
D
SOI-STJ Chip design pattern
LN2 termp.
Room temp.
SOI wafer I-V curve @Tsukuba
Vgs [V]
Ids [A]
Vds=0.2 V

## Slide 19: Summary
We are developing a detector to measure single photon energy with <2% resolution for .
  Our choices are Hf-STJ and Nb/Al-STJ array with grating.
We’ve confirmed to Hf-HfOx-Hf structure is established.
  Much improvement in leakage current is required.
  Improvement in I-V curve evaluation for Hf-STJ is more important.
Development of readout electronics for Nb/Al-STJ is underway.
  Aiming to measure a single photon of visible light at the first milestone.
  Several ultra low temperature amplifier candidates are under development.
    HEMT/SOI/SOI-STJ