## Slide 1
University of TsukubaA, KEKB, JAXAC, RIKEND, Okayama UniversityE, University of FukuiF, Kindai UniversityG, FermilabH
◯Kota KASAHARAA, S.H. KIMA, Y. TAKEUCHIA, R. SENZAKIA, K. NAGATAA, T. OKUDAIRAA, M. KANAMARUA, T. ICHIMURAA, K. MORIUCHIA, K. KIUCHIA, Y. ARAIB, M. HAZUMIB, T. WADAC, S. MATSUURAC, S. MIMAD, H. ISHINOE, Y.KATOF, T. YOSHIDAG, E. RAMBERGH, M. KOZLOVSKYH, P. RUVINOVH, D. SEGRATSKOVH
Development of Superconducting Tunnel Junction(STJ) Photon Detector on SOI Preamplifier Board to Search for Radiative Decays of Cosmic Neutrino Background
TIPP ‘14
International Conference on Technology and Instrumentation in Particle Physics ‘14
2nd-6th June 2014 in Amsterdam, the Netherlands.
Notes: I will be talking about Development of Superconducting tunnel junction on SOI Preamplifier board to search for Radiative decays of cosmic neutrino background.

## Slide 2: Outline
Motivation(Search for Neutrino Radiative Decay)
Introduction of Superconducting Tunnel Junction(STJ)
STJ on Silicon-on-Insulator Preamplifier-board(SOI-STJ)
Current Status on Development of SOI-STJ
Future Plan
Summary
TIPP ‘14
Notes: This is the outline of my talk.
Firstly, I will talk about the motivation and introduce the Superconducting Tunnel Junction.
Next, I will be talking about STJ on SOI Preamplifier board, we call SOI-STJ photon detector this device.
And current status on development of SOI-STJ.
Finally, I will make a summary and speak future plans.

## Slide 3: Motivation
TIPP ‘14
Lifetime of Neutrino(ν3 → ν1,2 + γ)
  1043 years in SM expectation, τ ⪎ O(1017) years in left-right symmetric model (SU(2)L⊗SU(2)R⊗U(1)B-L)
  Present limit is 3 x 1012 years (AKARI).
  We will use Cosmic Neutrino Background as neutrino source for Neutrino Decay.
Neutrino Mass Measurement
  We know that neutrinos have small masses but neutrino mass itself have not been measured.
  Aiming at measuring neutrino mass itself with Neutrino Radiative Decay.
Notes: We knew that neutrino has small mass by oscillation experiments, but neutrino mass has not determined.
Our motivation is determination of neutrino mass itself with neutrino Radiative decay.
But the lifetime of neutrino is too long in SM expectation. That’s calculated 10 to 43 years.
In some BSM, it is calculated shorted. For example, in Left Right Symmetric model it is calculated 10 to 17 years.
So we will use CNB as neutrino source for search for Radiative decay.

## Slide 4: Requirements for the Detector
TIPP ‘14
Neutrino Decay Energy Spectrum
Expected energy of photon in neutrino radiative decay is ( is assumed).
In this region, biggest background is zodiacal light.
To detect the neutrino radiative decay significantly in BG, We need the photon detector which has 2% energy resolution for 25meV(50um).
Red shift Effect
Zodiacal light
50 x 8 pixels
Nb/Al-STJ photon detector (50 x 8 pixels) with grating is a candidate for our experiment.
Aiming at improving the current experimental lower limit of the neutrino lifetime by 2 order in rocket experiment: (3)>O(1014 years)
80um
40um
Search region
Notes: Expected energy of photon in neutrino Radiative decay is 25meV assuming heaviest neutrino is 50meV.
In this region biggest background is Comic Infrared background, the energy spectrum is here.
The signal has sharp edge and the tail in lower energy due to red shift effect.
To detect the neutrino Radiative decay in too much background CIB, we need to photon detector which has 2% energy resolution for the edge.
We plan to use combined Nb/Al-STJ pixels with grating. And aim at improving the current experimental lower limit of the neutrino lifetime by 2order
In a 5minute measurement with this photon detector.
Anyway, the Nb/Al-STJ is required far infrared single photon detection.

## Slide 5: STJ Photon Detector
TIPP ‘14
STJ is a Josephson Junction composed of Superconductor / Insulator / Superconductor.
Photon absorbed in superconducting layers of STJ (Josephson current suppressed by applying magnetic field).
Excites cooper pairs into quasi-particles according to absorbed energy.
Observe tunneling current due to quasi-particles through the insulator.
Principle of Operation
 | Si | Nb | Al
Tc[K] |  | 9.23 | 1.20
Δ[meV] | 1100 | 1.550 | 0.172
Number of Quasi-particles in Nb/Al-STJ
GAl : Trapping Gain In Al(~10)
E0 : Photon Energy
Δ : E-Gap in superconductor
For 25meV single photon
Nq = 10
= 95 e
What’s STJ(Superconducting Tunnel Junction) ??
500nm
Notes: STJ is a Josephson junction composed of Superconducting layer / Insulator layer/ Superconducting layer.
If the photon incident to the STJ, firstly, photon absorbed in superconducting layer of STJ, then, excites cooper pair into quasi particles according to absorbed photon energy, finally we can observe tunneling current due to quasi particles through the insulator.
The number of quasi-particles in Nb/Al-STJ is described by this equation. For absorbing 25meV single photon, the number of Quasi-particles are about 100e.

## Slide 6: Development of cold preamplifier
TIPP ‘14
Current Status of Development for Nb/Al-STJ(if you want to know detail, please see Yuji TAKEUCHI’s slides-Ⅱ.c Neutrino session 12:00-).
  We haven’t succeeded in detecting a far-infrared single photon yet due to readout noise(thermal noise and cable noise).
  Need to adopt the preamplifier operated at ultra-low-temperature.
  FD-SOI device was proved to operate at 4K by a JAXA/KEK group(AIPC 1185,286-289(2009)).
Silicon on Insulator(SOI)
  Processing LSI on SiO2 Insulator.
  Very Small Parasitic Capacitance due to separation between MOSFETs with insulator(Not Depletion).
    Low-Power to be Operated
    High Speed
    Good at Large Scale Integration
    Suppression of charge-up by high mobility carrier due to thin depletion layer(~50nm).
Apply this to readout of STJ signal !
Bulk-CMOS
SOI-CMOS
~50nm
~10um
Notes: But actually, we have not succeeded in detecting a far-infrared single photon yet due to readout noise.
To solve this problem, we need to adopt the preamp can be operated at ultra low temperature we think.
Silicon on Insulator device is a candidate to improve the Signal to noise ratio.
This technology is processing LSI on SiO2 Insulator, so SOI device has very small parasitic capacitance due to separation between MOSFETs or the other device with insulator.
Consequently, SOI device has low power to be operated, high speed, good at LSI, and suppression of charge-up by high mobility carrier due to thin depletion layer.
So we think to employ SOI device to readout of STJ signal.

## Slide 7: Requirement for the amplifier of STJ readout.
TIPP ‘14
Speed
  We think that integration region of charge is from 1uS to 4uS.
Low power consumption
  Typical cooling power of our refrigerator is 400uW.
Operation at ultra-low temperature
  Requirement for leakage current of Nb/Al-STJ is below 100pA.
  But practically, leakage current is caused by…
    Pinhole(constant) across the insulator
      We can expect to reduce it by using smaller one.
    Thermal excitation()
      -We need to make cooler to 800mK.
100pA@800mK
To 100pA
Ileak∝
1. Operating below 800mK
2. Power consumption in the amplifier is as low as possible.
3. amplifying up to 1MHz is sufficiently good to apply.
Temperature Dependence of Dark Current with Nb/Al-STJ(10000um2)
3He sorption is better
Notes: Requirement of the amplifier for STJ readout are divided following three components.
One of these, the amplifier should be operated at ultra low temperature.
The requirements for leakage current of Nb/Al-STJ is below 100pA. But practically, leakage current is caused by pinhole and thermal excitation.
We can reduce the pinhole using smaller junction size. But to reduce thermal excitation, we need to make cooler to 800mK.
The other component is that the amplifier can be operated rather low power consumption.
Typical cooling power of our refrigerator is 400uW. So the power consumption in the amplifier is also below 400uW.
Another one, high speed.
Width of STJ signal is about 1.5uS. The amplifier can amplify up to around 2MHz.

## Slide 8
Via(W) for Electrical Contact
What’s SOI-STJ ?
Processing Nb/Al-STJ on SOI preamplifier board directly.
  Expected good signal-to-noise ratio.
  Potential for expansion of multi-pixel device
TIPP ‘14
The first prototype of SOI-STJ
We tested the Nb/Al-STJ on SOI wafer which only include MOSFET to answer the following questions:
SOI-FET has no damage by processing STJ ?
Quality of Nb/Al-STJ can be good on SOI wafer ?
SOI-STJ
STJ
Capacitance
FET
700 um
640 um
SOI-STJ
Circuit
STJ
C
Gate
Drain
Source
Notes: I will be talking about What is SOI-STJ. This is a main content in my presentation.
SOI-STJ is combined the STJ photon detector with ultra low temperature preamplifier.
Actually, we process Nb/Al-STJ on SOI preamplifier board directly.
So we expect that SOI-STJ improve the Signal to noise ratio, and has possibilities for expansion of Multi pixel device.
As a first step, we try to fabricate the Nb/Al-STJ on Soi wafer processing only MOSFET. In fact this is a schematic of this prototype.
And picture after processing Nb/Al-STJ is this. Around here, Via is for electrical contact between the STJ and SOI preamplifier.
Our purpose of Evaluation for this prototype are
SOIFET has no damage ?
Quality of Nb/Al-STJ can be good on SOI wafer?

## Slide 9: Nb/Al-STJ can be processed on SOI wafer perfectly ?
TIPP ‘14
Refrigerator
1K Ohm
Readout circuit
2 mV /DIV.
1 mA /DIV.
500uV /DIV.
10 nA /DIV.
2mV /DIV.
50uA /DIV.
2 mV /DIV.
1 mA /DIV.
Ic～200uA
ZOOM IN
ZOOM IN
Nb/Al-STJ
We have measured I-V curve of Nb/Al-STJ (50um x 50um junction) processed on SOI wafer
At 700mK with dilution refrigerator.
We have seen the distinctive I-V curve of Josephson Junction !
Leak current @ 0.5mV is 6nA. It’s almost same as our best record of normal Nb/Al-STJ(100um x 100um) 10nA.
Quality Factor (Rdynamic/Rnormal)
On Si wafer : 5 x 105
On SOI wafer : 3 x 105
B=35 G
Notes: Firstly, we have measured I-V curve of Nb/Al-STJ processed on SOI wafer at 700mK with dilution.
we have seen the distinctive I-V curve of Josephson Junction.
Left plot is before applying the magnetic fields, right one is during applying the magnetic fields.
This device has the leakage current 6nA. This value is almost same our best record.
We can confirm tha Nb/Al-STJ can be processed on SOI wafer perfectly.

## Slide 10
TIPP ‘14
1. SOIFET can be operated below 800mK ?
NMOS(690-750mK)
Width = 1000um
Length = 1um
PMOS(750mK)
Width = 1000um
Length = 1um
NMOS
PMOS
Excellent performance !!
I-V curve of SOIFETs after processing Nb/Al-STJ.
Both of NMOS and PMOS could be operated below 800mK.
Trans-conductance “gm” was not varied drastically for each temperature at operation voltage(0.2V).
  We can use these as Preamplifier for STJ signal
at ultra-low temperature.
Ids [A]
-Ids [A]
Vgs[V]
Vgs[V]
gm [S]
gm [S]
Vds[V]
Vds[V]
Notes: We have checked I-V curve of SOIFET can be operated below 800mK.
These plot are I-V curve of SOIFETs after processing Nb/Al-STJ. Left plot is described nmos at 690mK, right one is for pmos at 750mK.
And these plots are described the voltage between drain and source vs the trans-conductance which proportional to the GAIN. Tans conductance was not varied drastically for each temperature.
That’s Excellent performance.

## Slide 11
TIPP ‘14
2. Power Consumption of SOIFET
Width = 1.42um
Length = 0.42um
@830mK
Vgs=0.8V
Vgs=0.9V
Vgs=1.0V
Bias voltage of SOI-FET in saturation region(red line at right figure) : 0.5 V
Current (Ids)of FET in saturation region at Vgs = 0.8V : 0.09 uA
Power consumption = 0.5 V ×0.09 uA
= 45 nW/FET in case of W/L=1.42um/0.42um
We need to calculate the thermal model in rocket experiment firstly.
We investigate the possibility of using it around the threshold voltage(lower power and higher gain are expected, but lower speed).
Notes: In this slide, we will suggest the power consumption of SOIFET.
Typical Bias voltage of SOIFET in region of Saturation is 1.8V.
And current of FET in region of Saturation at Vgs = 0.8V at 830mK . 0.09uA in this FET.
So we can calculate power to be operated equal to 0.16uW.
This value means multipixels composed 50 times 8 equal t 65uW.
So the power consumption of SOIFET satisfies our requirement !!
回路設計はまだしていないが、実際にロケット実験の実験系でのthermal modelを調べる必要がある。また、SOIの消費電力に対してNb/Al-STJの温度上昇も評価するべき。
バッファ回路やデジタル信号処理等どこまでSOI上で行うかとか調べる必要がある。
もしも発熱が問題になった場合は、FETをthreshold voltage付近で用いることも考えている。

## Slide 12
TIPP ‘14
3. Frequency dependence of SOIFET
r0
500uV /DIV.
1uS /DIV.
Signal
(1.5uS)
We observed the signal of Nb/Al-STJ processed on SOI board to 465nm laser pulse.
The width of the signal of Nb/Al-STJ is about 1 uS.
Our requirement for the amplifier is operation above 1 MHz in order to best S/N.
We Checked the behavior of SOI-FET as the preamplifier.
But we can not measure the gain at high frequency due to too large output impedance.
Set GAIN 8.7 @ LHe temp.
Notes: And we also observed the signal of Nb/Al-STJ processed on SOI pc board to 465nm laser pulse. From this, the width of the signal of Nb/Al-STJ
Is about 1.5uS.
We want to distinguish photon by photon for the photon counting in order to best the single to noise ratio.
So our requirement for the amplifier is operation above few MHz.
We need to check it. We checked the frequency dependence of SOIFET as the preamplifier with left schematic.

## Slide 13
TIPP ‘14
Future Plans
We are updating the SOI-STJ schematic for amplification of the Nb/Al-STJ signal.
  Replace the resistance to SOI-FET that we use as current source.
  Employ the Feedback between drain and gate to apply stable bias voltage
  Add the follower to reduce the Output Impedance.
Designed the ratio (W/L) so that the operation power is below 120uW.
In current status, We will measure the response to laser pulse with renewed SOI-STJ soon.
Notes: So we are updating the SOI-STJ schematic for amplification of the Nb/Al-STJ signal.
The difference between renewed and old one as follows.
one of these, we replaced the resistance to SOIFET that we use as current source.
The other one , employ the feedback between drain and gate to apply stable bias voltage.
Another one, we added the follower to reduce the output impedance.
In current Status, we will measure the response to laser pulse with renewed SOI-STJ soon.

## Slide 14
TIPP ‘14
Summary
SOI-STJ photon detector is under development for detecting far-infrared single photon for neutrino mass measurement.
We processed Nb/Al-STJ on SOI-FET board and confirmed the following performances.
  SOIFET has no damage by processing Nb/Al-STJ.
  Nb/Al-STJ is also operated normally.
  SOIFET can be operated below 800mK.
  *For your information, we also confirmed that SOIFET can be operated at 100mK.
Notes: Finally, we make a brief summary. That is all.

## Slide 15
TIPP ‘14
Back Up Slides

## Slide 16
TIPP ‘14
How can we check the performance of STJ??
Josephson Junction has distinctive I-V curve.
The leakage current at operation voltage(0.5mV) deteriorate of energy resolution.
  To discriminate signal from leakage current fluctuation, the leak current of the Nb/Al-STJ should be below 100pA(Our best sample has 10nA).
Blue region
Cooper pairs can go through directly the other electrode due to larger bias than gap energy.
Red region
The cooper pair tunneling Current is seen at V = 0V.
This is suppressed during applying Magnetic fields.
The I-V curve of Josephson junction

## Slide 17
TIPP ‘14
Trapping Gain
STJ Back tunneling
When quasi-particle that is created around insulator go through the insulator layer, then quasi-particle break cooper pair in opposite layer and make cooper pair with one of quasi-particles.

## Slide 18
TIPP ‘14
I-V curve of SOIFET @ ultra-low temperature
SOI-FET
Gate
Metal
Channel
Width(w)
Channel
Length(L)
w=1.42um,
l = 0.4um
@830mK
w=1um×4, l = 1um
@830mK
w=10um×4, l = 1um
@830mK
We can see the I-V curve for each W/L value.

## Slide 19
TIPP ‘14
Requirement for leakage current of Nb/Al-STJ
When far-infrared single photon incident to the Nb/Al-STJ, Number of created quasi-particle in STJ is around 100e(assuming Gal = 10).
We want to separate between the signal and pedestal (3 sigma). so requirement for fluctuation of the leakage current of STJ is follow.
If we assume that integration width is 1.5 uS, our requirement for the leakage current is…

## Slide 20
TIPP ‘14
Temperature Dependence of I-V curve
Threshold voltage is changed. But the other properties are almost unchanged.

## Slide 21
TIPP ‘14
Energy resolution of Nb/Al-STJ
Energy Resolution of STJ
F : Fano Factor
Δ :Gap energy
E : Photon energy
= 22%
For 25meV single photon

## Slide 22
TIPP ‘14