## Slide 1: Development of Far-infrared Spectrophotometers based on Superconducting Tunnel Junction for the Cosmic Background Neutrino Decay (COBAND) Experiment
Sep 17-19, 2016 / Epochal Tsukuba
Yuji Takeuchi (Univ. of Tsukuba)
on behalf of COBAND Collaboration

## Slide 2: COBAND (COsmic BAckground Neutrino Decay)
Heavier neutrinos (2, 3) are not stable
  Neutrino can decay through the loop diagrams
However, the lifetime is expected to be much longer than the age of the universe
We search for neutrino decay using Cosmic Background Neutrino (CB) as the neutrino source

## Slide 3: COBAND Collaboration Members (As of Sep. 2016)
Shin-Hong Kim, Yuji Takeuchi, Kenichi Takemasa, Kazuki Nagata, Kota Kasahara, Shunsuke Yagi, Rena Wakasa, Yoichi Otsuka (Univ. of Tsukuba),
Hirokazu Ikeda, Takehiko Wada, Koichi Nagase (JAXA/ISAS),
Shuji Matsuura (Kwansei gakuin Univ),
Yasuo Arai, Ikuo Kurachi, Masashi Hazumi (KEK),
Takuo Yoshida，Chisa Asano，Takahiro Nakamura (Univ. of Fukui),
Satoshi Mima, Kenji Kiuchi (RIKEN),
H.Ishino, A.Kibayashi (Okayama Univ.),
Yukihiro Kato (Kindai University),
Go Fujii, Shigetomo Shiki, Masahiro Ukibe, Masataka Ohkubo (AIST),
Shoji Kawahito (Shizuoka Univ.),
Erik Ramberg, Paul Rubinov, Dmitri Sergatskov (Fermilab)，
Soo-Bong Kim (Seoul National University)
3

## Slide 4: Motivation of -decay search in CB
If we observed the neutrino radiative decay at the lifetime much shorter than the SM expectation, it would be
Physics beyond the Standard Model
Direct detection of CB
Determination of the neutrino mass
Aiming at a sensitivity to 3 lifetime in
Standard Model expectation:
Experimental lower limit:
Left-Right symmetric model predicts for - mixing angle
3 Lifetime

## Slide 5: Expected photon wavelength spectrum from CB decays
dN/d(a.u.)
[m]
100
500
10
Red Shift effect
Sharp Edge with 1.9K smearing
E [meV]
100
50
20
10
5
 distribution in
No other source has such a sharp edge structure!!
If assume
,
from neutrino oscillation measurements

## Slide 6: CB radiative decay and Backgrounds
at λ＝50μm
Zodiacal Emission
~8MJy/sr
Cosmic Infrared Background (CIB)
~0.1~0.5MJy/sr
~0.5MJy/sr
s
~25kJy/sr
Expected spectrum
for
CMB
ZE
ZL
ISD
SL
DGL
CB decay
wavelength [m]
E [meV]
Surface brightness I [MJy/sr]
AKARI
COBE
CB decay
excluded

## Slide 7
Proposal for COBAND Rocket Experiment
  JAXA sounding rocket S-520
    http://www.isas.jaxa.jp/e/enterp/rockets/sounding/s520.shtml
    Diameter: 520mm
    Payload: 100kg
    Altitude: 300km
Aiming at a sensitivity to 𝜈 lifetime for
200-sec measurement at altitude of 200~300km
    Telescope with diameter of 15cm and focal length of 1m
    All optics (mirrors, filters, shutters and grating) will be cooled at ~1.8K

## Slide 8
Proposal for COBAND Rocket Experiment
Nb/Al-STJ array
8 rows
50 columns
  At the focal point, diffraction grating covering =40-80m (16-31meV) and array of photo-detector pixels of 50(in wavelength distribution) x 8(in spatial distribution) are placed
Each pixel counts FIR photons in ()
Sensitive area of 100mx100m for each pixel (100rad x 100rad in viewing angle)

## Slide 9: Sensitivity to neutrino decay
Parameters in the rocket experiment simulation are assumed.
Can set lower limit on 3 lifetime at 4-6  1014 yrs if no neutrino decay
If 3 lifetime were 2  1014 yrs, the signal significance would be at 5 level
Current lower limit (S.H.Kim 2012)
x100 improvement!

## Slide 10: Requirements for the photo-detector in the rocket experiment
CMB
ISD
SL
DGL
CB decay
wavelength [m]
Surface brightness I [MJy/sr]
Zodiacal Emission
Zodiacal Light
Integrated flux from galaxy counts
Zodiacal Emission
=8MJy/sr
1.1×10-17 W / 8pix
@ =50m
CIB summary from Matsuura et al.(2011)
Neutrino Decay
=25kJy/sr
3.3×10-20 W / 8pix
@ =50m

## Slide 11: Noise Equivalent Power (NEP) Requirements for the photo-detector
Neutrino decay (, yrs): =25kJy/sr @ =50m
Zodiacal emission: =8MJy/sr @ =50m
Shot noise in PZE integrated over an interval t
  Fluctuation in number of photons with energy :
 >200sec
 for 1pix

## Slide 12: Existing FIR photo-detectors
Detectors | (μm) | Operation Temp. | NEP (W/Hz1/2) | 
Monolithic Ge:Ga | 50-110 | 2.2K | ～10-17 | Akari-FIS
Stressed Ge:Ga | 60-210 | 0.3K | ~0.9×10-17 | Herschel-PACS
We are trying to achieve by using
Superconducting tunneling junction detector
FIR single-photon counting technique
Need more than 2 orders improvement from existing photoconductor-based detectors

## Slide 13: Superconducting Tunnel Junction (STJ) Detector
A constant bias voltage (|V|<2) is applied across the junction.
A photon absorbed in the superconductor breaks Cooper pairs and creates tunneling current of quasi-particles proportional to the deposited photon energy.
Superconductor / Insulator /Superconductor Josephson junction device
Δ: Superconducting gap energy
2
E
Ns(E)
Superconductor
Superconductor
Insulator
Insulator
Superconductor
100m
300nm
Much lower gap energy (Δ) than FIR photon  Can detect FIR photon
Faster response (~s)  Suitable for single-photon counting

## Slide 14
STJ energy resolution
Signal = Number of quasi-particles :
Resolution = Statistical fluctuation in number of quasi-particles
Δ: Superconducting gap energy
F: Fano factor
E: Photon energy
G: Back-tunneling gain
Nb/Al-STJ
Well-established
Δ~0.6meV by the proximity effect from Al
Operation temperature <400mK
Back-tunnelling gain G~10
Nq.p.=25meV/1.7Δ10~ 250 E/E~0.1 for E=25meV
25meV single-photon detection is feasible in principle
Notes: 14

## Slide 15: STJ response to pulsed laser
100uV/DIV
4us/DIV
Nb/Al-STJ response to pulsed laser (465nm)
CRAVITY Nb/Al-STJ 100m sq.
VIS laser through optical fiber
STJ
V
Refrig.
V0
R0=1M
Nb/Al-STJ has ~1s response time.
We can improve NEP by photon counting in 1s integration time
  However we need faster readout system than f>1MHz
T~300mK

## Slide 16: Nb/Al-STJ development at CRAVITY
Temperature(K)
0.3
0.4
0.5
0.6
0.7
Leakage
100pA
1nA
10nA
100nA
Ileak~200pA for 50m sq. STJ, and achieved 50pA for 20m sq.
0.1nA
The shot noise from the leak current in case of ileak=50pA, =0.6meV and G=10
500pA/DIV
0.2mV/DIV
I
V
T~300mK
w/ B field
(Without photon counting)
50m sq. Nb/Al-STJ fabricated at CRAVITY
Photon counting will yield improvement by more than one order.

## Slide 17: 100x100m2 Nb/Al-STJ response to 465nm pulsed laser
Laser pulse trigger
We observed NIR-VIS laser pulse at few-photon level with a charge-sensitive amplifier placed at the room temperature.
Due to the readout noise, a FIR single-photon detection is not achieved yet.
Need ultra-low noise readout system for STJ signal
Considering a cryogenic pre-amplifier placed close to STJ
2V/DIV
40μs/DIV
Response is consistent to 10-photon detection in STJ
465nm laser through optical fiber
STJ
10M
T~350m
(3He sorption)
Charge sensitive
pre-amp.
shaper amp.

## Slide 18: FD-SOI-MOSFET at Cryogenic temperature
FD-SOI : Fully Depleted – Sillicon On Insulator
  Very thin channel layer in MOSFET
    No floating body effect caused by charge accumulation in the body
  FD-SOI-MOSFET is reported to work at 4K
Channel W
Channel L
    JAXA/ISIS AIPC 1185,286-289(2009)
    J Low Temp Phys 167, 602 (2012)

## Slide 19: FD-SOI MOSFET Id-Vg curve
Vgs (V)
Ids
0
0.5
1
1.5
2
1pA
1nA
1A
1mA
̶̶ ROOM
̶̶̶ 3K
n-MOS
p-MOS
̶̶ ROOM
̶̶̶ 3K
-Ids
1nA
1A
1mA
Vgs (V)
0
-0.5
-1
-1.5
-2
Id-Vg curve of W/L=10m/0.4m at |Vds|=1.8V
Both p-MOS and n-MOS show excellent performance at 3K (We confirmed they function down to 300mK).
Threshold shifts, sub-threshold current suppression and increase of the carrier mobility at low temperature.

## Slide 20: SOI prototype amplifier
8mV
150mV
10mV
100s
Ｖ
Amplifier stage
Buffer stage
T=350mK
Test pulse input through C=1nF capacitance at T=3K and 350mK
Power consumption ~100μW
INPUT
OUTPUT
1nF
Test pulse
We can compensate the effect of shifts in the thresholds by adjusting bias voltages.

## Slide 21: SOI charge-sensitive pre-amplifier development
STJ has comparably large capacitance: ~20pF for 20m sq. STJ.
  A low input impedance charge-sensitive amplifier is required for STJ single-photon signal readout.
STJ response time is ~1s.
We designed SOI op-amp which has >1MHz freq. response, and submitted to the next SOI MPW run. We’ll test the amplifier in this winter.
STJ
T~0.35K
Charge sensitive
pre-amp.
CSTJ
10M
I=I(V)

## Slide 22: Op-amp Circuit
Buffer stage
Telescopic cascode
Bias
telescopic cascode differential amplifier
Feedback C=2pF x R=5MOhm = 10s
Power consumption ~150W
Iref
Iref
Iref
Iref2
10/10
W(m)/L(m)
10/10
10/10
20/10
12/7
100/7
100/7
100/1
100/1
4/1
4/1
2/10
4/10
4/10
VDD=-VSS=1.5V
Iref>10A
Iref/5
Iref2/5
Iref2/5

## Slide 23: COBAND project
Nb/Al-STJ設計・開発
Japanese FY | 2016 | 2017 | 2018 | 2019 | 2020 | 2021
Setup design |  |  |  |  |  | 
STJ detector |  |  |  |  |  | 
Cryogenic electronics, readout circuit |  |  |  |  |  | 
Optics |  |  |  |  |  | 
Rocket-borne refrig. |  |  |  |  |  | 
Measurement Analysis |  |  |  |  |  | 
Rocket exp.
Nb/Al-STJ
（ＳＯＩ-ＳＴＪ） R&D
Production
R&D
Production
Analysis tool devel.
Analysis
Design
Production
Design
Production
Design
Design
Design
Hf-STJ R&D
Satellite exp.
Simulation
Rocket exp.

## Slide 24: Summary
We propose a sounding rocket experiment to search for neutrino radiative decay in cosmic neutrino background.
Requirements for the detector is a photo-detector with NEP~O(10-20) W/Hz
Nb/Al-STJ array with a diffractive for the experiment.
  Nb/Al-STJ fabricated at CRAVITY meets our requirements.
  FD-SOI readout is under development and almost ready for STJ signal amplification at cryogenic temperature.
Improvement of the neutrino lifetime lower limit up to O(1014yrs) is feasible for 200-sec measurement in a rocket-borne experiment with the detector.

## Slide 25: Backup

## Slide 26: Energy/Wavelength/Frequency

## Slide 27: Limit on LRSM parameters from TWIST
Manifest LRS 90%CL
Non-manifest LRS 90%CL

## Slide 28: Cosmic background neutrino (CB)
CMB
CB (=neutrino decoupling)
~1sec after the big bang

## Slide 29: Cosmic background neutrino (CB)
/generation
Density (cm-3)
The universe is filled with neutrinos.
However, they have not been detected yet!

## Slide 30: Neutrino Mass and Photon Energy
m3=50meV
m1=1meV
m2=8.7meV
Eγ =24meV
E =24.8meV
From neutrino oscillation
CMB (Plank+WP+highL) and BAO
50meV<<87meV
14~24meV
51~89m

## Slide 31: Sensitivity to neutrino decay
Parameters in the rocket experiment simulation
telescope dia.: 15cm
50-column (: 40m – 80 m)  8-row array
Viewing angle per single pixel: 100rad  100rad
Measurement time: 200 sec.
Photon detection efficiency: 100%
Can set lower limit on 3 lifetime at 4-6  1014 yrs if no neutrino decay
If 3 lifetime were 2  1014 yrs, the signal significance would be at 5 level

## Slide 32: STJ current-voltage curve
Optical signal readout
Apply a constant bias voltage (|V|<2) across the junction and collect tunneling current of quasi particles created by photons
Leak current causes background noise
Tunnel current of Cooper pairs (Josephson current) is suppressed by applying magnetic field
Leak current
B field
I-V curve with light illumination
Voltage
Current
Signal current

## Slide 33: STJ back-tunneling effect
Photon
Bi-layer fabricated with superconductors of different gaps Nb>Al to enhance quasi-particle density near the barrier
  Quasi-particle near the barrier can mediate multiple Cooper pairs
Nb/Al-STJ Nb(200nm)/Al(70nm)/AlOx/Al(70nm)/Nb(200nm)
Gain: ～10
Nb
Al
Nb
Al
Si wafer
Nb
Nb
Al/AlOx/Al

## Slide 34: SOI-STJ development
STJ layers are fabricated directly on a SOI pre-amplifier board and cooled down together with the STJ
Started test with Nb/Al-STJ on SOI with p-MOS and n-MOS FET
SOI
STJ
Nb
metal pad
STJ lower layer has electrical contact with SOI circuit through VIA
VIA
STJ
capacitor
FET
700 um
640 um
C
SOI-STJ2 circuit
D
S
G

## Slide 35: FD-SOI on which STJ is fabricated
Both nMOS and pMOS-FET in FD-SOI wafer on which a STJ is fabricated work fine at temperature down below 1K
Nb/Al-STJ fabricated at KEK on FD-SOI works fine
We are also developing SOI-STJ where STJ is fabricated at CRAVITY
B~150Gauss
2mV/DIV
1mA/DIV
I-V curve of a STJ fabricated at KEK on a FD-SOI wafer
nMOS-FET in FD-SOI wafer on which a STJ is fabricated at KEK
gate-source voltage (V)
drain-source current
0
0.2
0.4
0.6
0.8
-0.2
1pA
1nA
1A
1mA

## Slide 36: Detector NEP required for COBAND rocket experiment
What’s NEP (Noise Equivalent Power)?
The signal power yielding unit signal-to-noise ratio in unit bandwidth for a photo-detector
  F: Radiant flux(W)
  S/N: signal-to-noise ratio of the detector ouput
  Δf : detector bandwidth(Hz)
A smaller NEP means a more sensitive detector.
A detector with can detect
  1pW signal with S/N=1 after 0.5-sec integration.
  0.1pW signal with S/N=1 after 50-sec integration.

## Slide 37: Detector NEP required for COBAND rocket experiment
Viewing angle of the photo-detector
Telescope main mirror: D=15cm, F=1m
100m x 100m x 8 pixels
  Viewing angle : 8 x (100m/1m)2 = 8 x 10-8 sr
Hereafter on the assumption of 100% quantum efficiency at the photo-detector
 per one wavelength division after diffraction grating:
One wavelength division: (80m - 40m) / 50 = 0.8m
=c/50m-c/50.8m=94GHz @=50m

## Slide 38: Detector NEP required for COBAND rocket experiment
Neutrino decay (, yrs): =25kJy/sr @ =50m
Zodiacal emission: =8MJy/sr @ =50m
The fluctuation in the FZE integration over t interval
  The fluctuation in # of photons:
Conditions of the requirements on t and NEP
 t>40sec (1) t>200sec (2.2 per =0.8m)
 for 8 pix
 for 1pix

## Slide 39: Shot noise from a detector leak current
R [A/W] : Detector responsivity
iL [A] : leak current
The charge from the leakage after T [sec] integration: iLT [C]
The number of charge carriers after averaging: iLT/e
The fluctuation in the current:
The fluctuation in the measured incident power:
i.e.
For STJ: i.e.