## Slide 1: Run2b Silicon Detector - Overview
K. Hara
U of Tsukuba
Project Status
Sensor Prototype:
Delivery
Testing at U Tsukuba - covered by Takei
Limitations in QA
Charge Up Issue
Sensor QA Agreements
Schedule

## Slide 2: Project Status
Layer 0: 12 fold Axial
Layer 1: 6 fold Axial-Axial
Layer 2: 12 fold Axial-Stereo (1.2o)
Layer 3: 18 fold Axial-Stereo (1.2o)
Layer 4: 24 fold Axial-Stereo (1.2o)
Layer 5: 30 fold Axial–Axial
Layout:
changes：
no 90o; 2.5o → 1.2o
6-fold sym for SVT
simplified L1
1512 outer axial sensors
648 outer stereo sensors
144 inner axial sensors
36x2 axial-axial staves
54x2 axial-stereo staves
passed Lehman Review (Sept 24,02)
DOE Review (Dec 18) production phase

## Slide 3: Sensor Prototypes: Delivery (HPK)
60 Axial Prototypes delivered (July 12)
30 grade “A”+30 grade “B” – actually all are grade “A” quality
40 better sensors are shipped to FNAL
53 Stereo Prototypes delivered (by end Oct)
18 + 18 sensors are shipped to FNAL
stereo delivery delayed by 2-3 wks due to
a process component was replaced
unexpected flaw on the surface ( grade B)
- wafers had to be taken manually from the line
electrically they are as good as axial sensors
For the order of 60 sensors for each:

## Slide 4: Visual Inspection
stains on stereo sensors
Fig. 8a Example of scratches observed for stereo sensors. (Right) Enlarged view.
about 2/3 of Class B have similar stains
No correlation with defect strips
No obvious correlation with I-V stability

## Slide 5: Limitations in QA - Implant Opens
Implant open is seen clearly,while HPK could not find. In total 3 such channels out of 36 sensors tested.
Fig. 9 Illustration of measurement configurations of readout and intermediate implant opens. strips .measurements.
passivated for
hybrid mounting
HPK detects implant opens by:
DC scan
- inject V on DC and measure I
into bias-ring: insensitive for readout strips
AC scan
- inject V on AC and measure transient I into bias-ring: not very sensitive if the open is at far side

## Slide 6: Limitations for stereo sensors
probed end:
511 AC pads
hybrid region
12 or 19 un-probed strips:
no opening in B.R., AC pads
Should we open for probing?
Measurement takes time only for these strips…
Our decision:
“Accept un-probed strips (and readout opens) since the dead fraction is very small”.

## Slide 7: Charge Up
In early stage, we often got bad distributions (R/Ccp e.g.)
Rint, Cint are also affected
From the location and size of bad strips, we suspected vacuum tweezers caused charge-up on sensor surface
After eliminating use of tweezers, we saw a few channels which showed initially “charge-up” effect and recovered later.
- This could be due to bad probe contact, though…
Anyway, the sensors are after all free from charge up

## Slide 8: Charge Up – its mechanism
<100> is easier than <111> to charge up
due to lower transition potential for inversion layer creation
+ + + + + + +
+ + + + + + +
fixed oxide charges in SiO2
Si-SiO2 interface
<111>:
about one order abundant positive charges
attracts more electrons, which isolate P+ implants
stronger against external effects
- - - - - - - - - - - - -
n
+ + +
+ + +
- - - - -
p+
<100>:
fewer positive charges
fewer electrons attracted
less robust against external effects
may create P layer = inversion layer

## Slide 9: Charge Up – its recovery
Inversion layer should disappear by radiation
accumulation of positive charges in
oxide layer, which attract more electrons
(> extra holes in inversion layer)
Surface charges responsible for inversion layer creation should be taken away if there is a route
nearby Al strips are wire-bonded
substantial surface current (UV?)
Extra epoxy on surface might disturb this recovery
Fig. 10b Photo of the sensor used for charge-up study. Locations of wirebonds and epoxy are illustrated.
Prepared a sample heavily charged
apply epoxy (Araldite 2011)
some Al strips wire-bonded

## Slide 10: Charge-Up Recovery (w/o UV)
epoxy
Bias is kept on at 200V
initial: uniformly charge-up
T=0: right after epoxied and wire-bonded (epoxy not cured)
Recovery is recognized
in epoxied region
in region close to wire-bonds
wirebonds
please ignore isolated points (bad contact)

## Slide 11: Charge-Up Recovery (w/ UV)
Bias is kept on at 200V
UV on (Ib~175uA)
substantial recovery in the region close to wire-bonds
(see the time constant)
Unfortunately but…
probes wiped prior to T=63h measurement
sensor taken off the stage !
please ignore isolated points (bad contact)
charge-up disappeared…
we terminated data taking

## Slide 12: When the sensor is removed from the stage
completely recovered
additional charge-up!!!
Lessons:
wire-bonds remove surface charge
epoxy helps remove surface charge
surface charge in isolated area could remain at least for days
UV enhances charge removal (near wire-bonded region)
epoxy degrades isolation, the system becomes weak against external charges
avoid use of tweezers unless it is proven not to charge-up
please ignore isolated points (bad contact)

## Slide 13: Document of Testing Procedure+QA
QA document is under circulation:
http//hep-www.px.tsukuba.ac.jp/~hara/run2bsi/QA_sensor.doc

## Slide 15: Schedule
2003
2004
Schedule driven by
stave construction/testing of components (hybrids, miniPC, bus cables)
Installation should complete in 6 month shutdown
1st prototype stave available
testing prototype DAQ chain complete
prototypes
order
production delivery
pre-production go ahead
2005
stave installation complete
outer detector complete
ready for installation into ISL
prototype L0 complete
milestones
L2 milestones: incl. 30wk total contingency
Si
prototypes
L0 cable
production cable available
production delivery
L0 Si
order
workshop: Jan22
sensor sign-off mtg: beg. Feb