Paper to be integrated into Nov. 24th, 2002 Draft v0

Jan. 15th, 2003 Draft v0

# Characteristics of Prototypes Silicon Sensors for run2b

#### D.Bortoletto, G.Bolla, A.Canepa

Purdue University, Purdue, West La Fayette, Indiana

## Abstract

- Introduction

- Production of Prototypes Sensors

Four Outer Axial Sensors were delivered at Purdue University, Lot No. SWA61457, Serial No. 0060, 0061, 0063 and 0069. For these sensors, HPK reports no failures.

3. Electrical characteristics

The sensors have been tested at the P3MD Lab probe station for conformance with the specifications. The probe station consists of Alessi Microscope, Keithley 617 Programmable Electrometer, Keithley 237 High Voltage Source Measure Unit and QuadTech 7600 Precision LCR. In order to perform an appropriate testing of the irradiated sensors, a cooling system has been developed to lower the chuck temperature (up to T=10C); this system allows to extensively probe unbonded sensors. An environmental chamber has been used to measure the characteristics at temperature as low as T=–25C; in this case, the sensors have to be bonded to a wire card.

The measurements we have conducted are listed in Table 1.

| MEASUREMENTS |  |
| --- | --- |
| I-V curve | The total leakage current is measured as a function of the bias voltage up to 1000V. The guard ring is floating (floating guard). |
| Stability I-V | Non-IRRADIATED sensors: the total leakage current is measured every 30 min. The sensor is polarized at Vbias=200V between the measurements. IRRADIATED sensors: the total leakage current is measured with the sensors biased at Vbias=300V. The guard ring is floating in both cases. |
| C-V curve | The total capacitance is measured as a function of the bias voltage. The depletion voltage is extracted by a fit to the data taken at f=10kHz and VAC =1V |
| AC scan | The coupling capacitance is measured at f=100 Hz and VAC =1V. The sensor is biased at Bias Voltage Vbias=10V. |
| AC punch-through | The coupling oxide is polarized at V=100V and the leakage current is measured. The sensor is polarized at Bias Voltage Vbias=10V |
| # DC scan | Strip leakage current is measured on a sample of strips for the sensors showing high total leakage current. |
| Bias Resistance | The bias resistance is measured by applying a voltage ramp between the DC pad and the bias line. The sensor is biased at Vbias= 300V |
| Interstrip Capacitance | The capacitance between two neighbors strips is measured at f=100 Hz and VAC =1V. The sensor is biased at Bias Voltage Vbias= 300V |

Table 1. Measurements summary

3.1.1 I-V curves

As required, HPK performed I-V measurement on all sensors that were delivered; data referring to the sensors mentioned above are listed in Table 2.

| HPK MEAUREMENT | ### Total Leakage Current at Bias Voltage V=1000V |
| --- | --- |
| 0060 | 126.99 nA |
| 0061 | 114.77 nA |
| 0063 | 123.27 nA |
| 0069 | 130.93 nA |

Table 2. Hamamatsu results. I-V measurements

The following plot (Fig.1) shows the I-V curves measured at Purdue: the sensor is biased from the backside and the guard ring is floating.

Fig.1 Purdue results. Leakage current versus bias voltage

We notice that sensors No.60 and No.61 have a total leakage current value that satisfies the production specification (ILeak < 50 nA/cm2 at Vbias=500V):

- sensor No. 60 Total Leakage Current ILeak = 3.6 nA/cm2 at Vbias=500V

- sensor No. 61 Total Leakage Current ILeak= 12.6 nA/cm2 at Vbias=500V

Sensors No.69 and No.63 show low breakdown voltage (No.69, Vbd= 180V; No.63, Vbd= 480V). The I-V measurement has been repeated by grounding the guard ring: no sensor breakdown voltage has been detected on sensor No.69; this means that the contribution driving to the breakdown is due to the edge of the sensor. On the other hand, the I-V curve for sensor No.63 does not change; further tests reveal that individual strip current shows the same behavior as the total current (the measurements is performed on 1/50 of the readout strips).

3.2 Stability I-V curves

The I-V stability test consists of a sequence of I-V measurements performed on a given sensor. The measurements are repeated every 30 min at room temperature; the bias voltage is set at Vbias=200V between the measurements. Fig.2 shows data for sensor No.61.

Fig.2 Purdue results. Leakage current versus time

No deviations from the original curve are developed as a function of time as the measurements are repeated on the given sensor. The test is repeated after 24h and the shape of the I-V characteristics is more stable than after the first measurements. This may be due either to different experimental conditions or to a stabilization of the surface effect.

3.3 C-V curves

The total sensor capacitance is measured as a function of the bias voltage. Fig.3 shows data obtained at an LCR frequency f=100kHz and VAC=1V. The depletion voltage is calculated as the intercept of the two straight lines in the 1/C2-Vbias plot. Most of the sensors satisfy the specification required for the depletion voltage (120V<Vdep<250V); sensor No.60 is closed to the lower limit of our selection. The reason of the discrepancy might be the choice of the fit. It should be noted that HPK results are always larger than the measured values. The Purdue results are consistent with the measurements performed at Tsukuba.

Fig.3 Purdue results. Total capacitance versus bias voltage

3.4.1 AC scan

The evaluation of the sensor performance is conducted in two steps on 1/10 of the readout strips.

- The coupling capacitance is measured by using the LCR in Cs-Rs mode at a frequency f=100Hz and VAC=1V. The probes are located at the AC pad and at the bias line. The sensor is polarized at Vbias=10V.

Fig.4 Purdue results. Coupling capacitance versus strip number

In Fig.4 the coupling capacitance is plotted as a function of the strip number and frequency for the three sensors that will be irradiated. The value of capacitance meets the requirements (CC > 120 pF) and the dependence on the frequency follows our expectations. The average values of CC are:

- sensor No.60, CC=134.43+/-0.48 pF

- sensor No.63, CC=125.32+/-1.99 pF

- sensor No.69, CC=126.49+/-0.97 pF.

Two channels in sensor No.63 show irregular values that might be due to a metal open; for this reason, strips have been inspected under the microscope but no evidence of open in the metal strip has been spotted. Furthermore, HPK did not detect both failures.

- The punch-through current is measured by using the electrometer: the coupling capacitance is polarized at Vpad=100V by probing between the AC pad and the bias line. The sensor is biased at Vbias=10V.

Fig.5 Purdue results. Punch-through current versus strip number

For example, Fig.5 shows that no pin-holes have been detected on sensor No.60 (in presence of a pin-hole the value would rise to A).

3.4.2 Bias Resistance

To evaluate the bias resistance, the poly-silicon is directly probed between the DC pad and the bias line while the sensor is polarized at Vbias > Vdep. This test allows spot any surface charge: if static charge is present on the sensor, the value of the resistance drops drastically since strips are shorted. These phenomena are detected on sensor No.69 and Fig.6 shows the results of the measurements. The static charge can be partially removed by cleaning the detector and locating it into a dry box.

Fig.6 Purdue results. Bias Resistance versus strip number

In the region that is NOT affected by surface charge the resistance is as expected (Rbias=1.5+/-0.5 MOhm); we safely claim the specifications are satisfied since the surface charge decays in time and can be totally removed by grounding the strips (see Tsukuba results).

3.7 Interstrip capacitance

The interstrip capacitance is measured between two neighbor strips at an LCR frequency f=1MHz and VAC=1V; the sensor is biased at Vbias =300V  (LCR mode: Cp-Rp).

The upcoming figure (Fig.7) shows the interstrip capacitance as a function of the strip number; the specification (Cint<1.2 pF/cm) is satisfied for all tested sensors, the average values being:

- sensor No.60, Cint=3.46+/-1.68 pF,

- sensor No.63, Cint=3.17+/-0.01 pF

- sensor No.69, Cint=3.53+/-0.18 pF.

We notice a trend in the region U100-U200 for both sensor No.60 and No.69. It should be recalled that sensor No.69 is strongly affected by surface charge (see 3.4.2 and appendix); nevertheless the charge density is detected on the other side of the sensor (region D200).

Fig.7 Purdue results. Interstrip capacitance versus strip number

Paper to be integrated into Nov. 24th, 2002 Draft v0

6.Performance of Neutron-Irradiated Sensors

Sensors were irradiated at a nominal dose =1.4*1014 1MeV eq-n at MNRC Irradiation Facility (UC Davis, California, September 27th 2002) and have been tested in Purdue.  As already mentioned, each sensor has been studied at the probe station, provided with a cooled-down chuck.  Thanks to this set-up, the sensor is biased from the backside and can be extensively probe. For further studies, each sensor has been bonded to a wire card and tested inside an environmental chamber (the operative temperature was set to T=-25C). The wire card is a frame with six ceramic pads provided with coaxial cables (one of these pads will be bonded to the n+ open to polarize the sensor from the top). Due to the card configuration a given number of strips only can be tested. We started testing sensors on October 23th.

6.1 Total Leakage Current

Sensors were always kept at a low temperature (T=-7C) since the irradiation was performed except when tested. In order to cross-check the dose by measuring the value of the current at T=20C sensor No.69 was annealed 4min at T=80min. Fig.8 shows he I-V characteristics for the mentioned detector. The early breakdown is due to a problem in the experimental set-up, later solved.

Fig.8 Purdue results. Leakage current versus bias voltage

The estimated flunce is lower than the nominal value: calc 1.1*1014 1MeV eq-n. This difference can be explained by taking into account the testing time (2500 min at room temperature); the fluence may also differ from the nominal value.

After irradiation the performances of all sensors have improved as shown in Fig.9: sensor No.63 and sensor No.69 having low breakdown voltage before irradiation show now a stable behavior: the charge depletion starts from the n+-p junction (inverted bulk) so that there are no region at high electric field.

Measurements are performed at T=13C and data are normalized to T=20C.

Fig.9 Purdue results. Leakage current versus bias voltage

A difference in the values of the current is observed although they were all irradiated at the same dose. Each sensor has been kept at room temperature for different periods of time (sensor 60, 1700 min; sensor 63, 1200 min; sensor 69, 2800 min); sensor No.69 were also annealed, as already mentioned.

The stability for sensor No.60 has been checked biasing the sensor while in the environmental chamber.

Fig.10 shows that no deviation from the given value takes place on the time scale of 1 h.

Fig.10 Purdue results. Leakage current versus time

6.2 Full Depletion Voltage

The measurement is strongly affected by high leakage current and therefore is performed in the environmental chamber at T=-25C. The LCR is set to f=1kHz and VAC=1V. The depletion voltage is again calculated as the intercept of the two straight lines in the 1/C2 –Vbias plot and normalized to the nominal frequency and temperature f=10kHz and T=20C (ROSE/TN/2000-10).

Fig.11 Purdue results. Total capacitance versus bias voltage

As an example, Fig.11 presents the behavior of the total capacitance for sensor No.60 measured 56 days after the irradiation. The normalized value is Vdep=128V. This value is lower than predicted; there are two reasons that may explain the discrepancy: the initial depletion voltage is not Vdep=150V (Vdep=117V, see 3.3) and the dose differs from the nominal value.

AC Scan

The AC scan is repeated on the irradiated sensors; it can be seen that the coupling capacitance still satisfies the specifications (Fig.12). The averages values measured on sensor No.60 and No.69 are CC=136.34+/-1.17 pF and CC=134.47+/-5.34 pF, both comparable to the pre-irradiation data. A stability test for coupling capacitance has been performed on a sample of strips. Fig.13 shows the time dependence of the punch-though current for a given readout strip: no spikes are reported and the integrity of the capacitance is verified up to Vpad=100V.

The bias resistance is measured to verify the strip isolation after irradiation. Since the current increases in time during the measurement, the value of resistance that can be measured is expected to be lower than the physical one (Fig.14) but again it satisfies the specifications. In Fig.15 the comparison between the values of the resistance before and after irradiation can be found. We notice that the drop due to static charge is not present anymore and the behavior follows our expectations. This can be explained by assuming both charge decay and irradiation effects (the charge trapped in the oxide changes due to irradiation). The performances of the detector improve after irradiation.

Fig.12 Purdue results. Coupling capacitance versus strip number

Fig.13 Purdue results. Punch-through current versus time

Fig.14 Purdue results. Bias resistance versus bias voltage

Fig.15 Purdue results. Bias resistance versus strip number

6.4 Interstrip Capacitance

The interstrip capacitance shows a time dependent behavior that may be correlated to the time decay of surface charge; this induces irregularities, such as high value of the capacitance at Vbias<500V (see appendix), that disappear as testing goes on. Therefore in Fig.16 and Fig.17 we show the measurements after irradiation both versus voltage and strip number and compare them with the results of measurements before irradiation.

Fig.16 Purdue results. Interstrip capacitance versus strip number

Fig. 17 Purdue results. Interstrip capacitance bias voltage

This measurement is performed at temperature T=14C to decrease the effect due to the large current after irradiation. It can be seen that the value is comparable with the pre-irradiation one.

Conclusions

In this work we have firstly verified the conformance to the specifications of four Hamamatsu sensors. Secondly, we have studied the effects of neutron irradiation. We observe a good agreement between the requirements and the results of measurements on un-irradiated detectors. Two sensors having low breakdown voltage as un-irradiated show stable behavior after irradiation; the depletion voltage for the inverted bulk has been measured to be Vdep=130V. The characterizing quantities (such as interstrip capacitance, bias resistance, coupling capacitance) are not affected by irradiation and therefore we conclude that the sensors are radiation hard up to =1.4*1014 1MeV eq-n.

Two channels appear to have open metal that was not detected by the vendor due to a low sensitivity in finding such defects when located on readout strip. A charge up tendency has been observed: the main effects are an increase of the interstrip capacitance at Vbias<500V and a decrease in the bias resistance. This feature is proved to decrease as the readout strip are grounded and decays in time. Therefore we conclude that Hamamatsu sensors comply our specifications for Run2B.